NATIONAL LEGAL METROLOGY EVALUATION CENTRE

OIML R 76 NON-AUTOMATIC WEIGHING INSTRUMENT TEST REPORT

Report No: LM/2026/00002 | Test ID: TEST-2026-0001 | Overall Result: INCOMPLETE

1. Instrument Under Evaluation

| Manufacturer | Avery Weigh-Tronix India Ltd. |
| --- | --- |
| Model / Serial No | WeighMaster Bench Scale / SN-2026-WM-98101 |
| Accuracy Class | Class III (Medium) |
| Max / Min Capacity | 30.000000 kg / 0.100000 kg |
| Scale Interval (e / d) | e = 0.010000 kg, d = 0.010000 kg |

2. Environmental & Testing Conditions

| Temperature / Humidity | 21.50 °C / 45.00 % |
| --- | --- |
| Atmospheric Pressure | 1012.50 hPa |
| Reference Mass Standards | E2 Mass Set (Cert: NPL-2026-CAL-101) |

3. Test Observations & Compliance Results

Test: Weighing Performance Test (Clause A.4.4) — Result: PASS

| # | Applied Load | Indication | Error (E) | MPE Limit | Status |
| --- | --- | --- | --- | --- | --- |
| 1 | 0.000000 kg | 0.000000 kg | 0.000000 kg | ±0.005000 kg | PASS |
| 2 | 5.000000 kg | 5.002000 kg | 0.002000 kg | ±0.005000 kg | PASS |
| 3 | 10.000000 kg | 10.003000 kg | 0.003000 kg | ±0.010000 kg | PASS |
| 4 | 20.000000 kg | 20.004000 kg | 0.004000 kg | ±0.010000 kg | PASS |
| 5 | 30.000000 kg | 30.005000 kg | 0.005000 kg | ±0.015000 kg | PASS |
| 6 | 20.000000 kg | 20.004000 kg | 0.004000 kg | ±0.010000 kg | PASS |
| 7 | 0.000000 kg | 0.000000 kg | 0.000000 kg | ±0.005000 kg | PASS |

Test: Repeatability Test (Clause A.4.10) — Result: PENDING

Test: Eccentric Loading Test (Clause A.4.7) — Result: PENDING

4. Verification & Approval

| Technician Signature | Priya Verma (Metrology Engineer) Date: 2026-09-19 |
| --- | --- |
| Reviewer Signature | Arun K. Patel (Senior Verification Officer) Date: 2026-09-19 |