NATIONAL LEGAL METROLOGY EVALUATION CENTRE

OIML R 76 NON-AUTOMATIC WEIGHING INSTRUMENT TEST REPORT

Report No: LM/2026/00016 | Test ID: TEST-2026-0017 | Overall Result: FAIL

1. Instrument Under Evaluation

| Manufacturer | Avery Weigh-Tronix India Ltd. |
| --- | --- |
| Model Name | WeighMaster Bench Scale |
| Serial Number | SN-2026-WM-98101 |
| Accuracy Class | Class III (Medium) |
| Capacity & Scale Interval | Max: 30.000000 kg, e: 0.010000 kg |

2. Environmental Conditions

| Temperature | 22.00 °C |
| --- | --- |
| Relative Humidity | 48.00 % |
| Atmospheric Pressure | 1013.25 hPa |

3. Test Observations & Evaluation

Procedure: Weighing Performance Test (Clause A.4.4) (Status: COMPLETED)

| Pt # | Applied (kg) | Indicated (kg) | Error (kg) | MPE (kg) | Verdict |
| --- | --- | --- | --- | --- | --- |
| 1 | 0 kg | 0 kg | 0 kg | ±0.005 kg | PASS |
| 2 | 500 kg | 500 kg | -500 kg | ±0.005 kg | FAIL |
| 3 | 1000 kg | 1000 kg | -1000 kg | ±0.005 kg | FAIL |
| 4 | 2000 kg | 2000 kg | -2000 kg | ±0.005 kg | FAIL |
| 5 | 3000 kg | 3000 kg | -3000 kg | ±0.005 kg | FAIL |

Procedure: Repeatability Test (Clause A.4.10) (Status: PENDING)

Procedure: Eccentric Loading Test (Clause A.4.7) (Status: PENDING)

4. Photographs & Attached Supporting Evidence

| Category | Original File Name | File Size | Uploaded By |
| --- | --- | --- | --- |
| INSTRUMENT PHOTO | nawi_test_photo.png | 0 KB | Dr. Rajesh Sharma (Lab Director) |
| NAMEPLATE PHOTO | nawi_test_photo.png | 0 KB | Dr. Rajesh Sharma (Lab Director) |

5. Official Metrological Verification & Digital Signature

| Digitally Signed By | Dr. Deepali Sharma (Chief Verification Reviewer) (Chief Metrological Verification Reviewer) |
| --- | --- |
| Certificate Serial / Reference | CERT-OIML-2026-IN-0017 [SIG-LM-2026-989D0F9B] |
| Signing Timestamp | 2026-09-30 15:57:37.120642 IST |
| Cryptographic SHA-256 Seal | c6efe17f97b9118376a695f70d5d8005c4e563824676203310ff5b648971b212 |
| Legal Metrology Declaration | I hereby certify under statutory Legal Metrology authority that this evaluation has been verified in strict compliance with OIML R 76-1:2006. |